## Slide 1
1. Apply factory calibration
CHLA = (CHLA_FLUORESCENCE470_RAW – DARK) * SCALE
2. Initial global range QC
Flag values outside -0.2 and 100
(QC flag = 4)
3. Despike
CHLA_BASELINE, SPIKES = despike(CHLA)
4. Dark correction
4.1 Keep profiles that meet depth criteria
Deeper than # dbar (user defined)
Exclude top 10 dbar
4.2 Find minimum of each profile
4.3 Calculate MAD
4.4 Flag outliers
4.5 Determine offset
Remove flagged values and calculate
5.2 Apply offset
CHLA_ADJUSTED = CHLA – Doffset
6. Apply NPQ correction
CHLA_ADJUSTED = NPQ correction(CHLA_ADJUSTED)
5.1 Apply offset
CHLA_FLUORESCENCE_ADJUSTED = CHLA – Doffset
KEY DIAGNOSTIC PLOTS
0.10
0.05
0
-0.05
-0.10
Profile number / Time
QC Flag = 1
QC Flag = 4 (High)
QC Flag = 4 (Low)
7. Final global range QC
Flag values outside 0 and 100
(QC flag = 4)
KEY DIAGNOSTIC PLOTS
Plot of Day vs Night CHLA_FLUORESCENCE_ADJUSTED
(Current one)